## Slide 1: PFS ICS infrastructure
Atsushi Shimono (PFS Project Office)
2017/03/09

## Slide 2: Agenda
PFS ICS infrastructure design review
  Overview of targets in this review
  Brief overview of PFS ICS design
  Network
  Link over fiber on telescope
  Storage and image data handling
  VM infrastructure
  Database server and its replication/backup
  ICS infrastructure status monitoring and defect detection
  Procedure on instrument exchange
  Hardware delivery to Subaru
Discussions on interfaces between Subaru and PFS
  Network connection to Subaru (data flow, V-LAN)
  Access control to/from PFS network
  System alert handling (items to STS)
  Engineering data archive (during on-site AIT, independent unit test)

## Slide 3: Overview of targets in this review
Network
Link over fiber on telescope
Storage and image data handling
VM infrastructure
Database server and its replication/backup
ICS infrastructure status monitoring and defect detection
Procedure on instrument exchange

## Slide 4: Items for review shown in ICS architecture diagram
Hardware
Pure-software
Individual
Integrated
Subsystem actor
Sequence (SpS/PFI)
Integrated control
Message Hub System (MHS)
Log/Status archive
Status viewer GUI
ICS Storage
Virtual Machine Infrastructure
PFS Network (PFS-LAN)
Backup infrastructure
TARGET AREA

## Slide 5
Status archiver
MHS
jsonp provider
SHM
Operational status viewer (browser)
database
DB to json
SQL
HTML
jsonp or CORS
summit
hilo
hilo outside-DMZ
DB
Summit master iSCSI storage
operation
backup
db server
DB
Hilo storage?
backup
db server
replication
Daily backup
Daily backup
External access
MHS
Status archive
(ICS)
operation
(on-site DRP)
Survey cluster?
Survey database
(On-site database backup/replication proposal)
Gen2
MHS
g2t
g2cam
m2t
IIC
MPS
MCS
MAC
MLP1
SAS
STS
Subaru
B/R/NCU1
ENU1
2
2
3
3
4
4
etc. …
DB
Storage
STARS
Hilo
On-site DRP
FPS
AGCC
TARGET AREA
TARGET AREA

## Slide 6: Brief overview of PFS ICS design

## Slide 7: Network
PFS has four IR detectors which make continuous data flow from devices to storage server(s), and its own camera system for auto guide. PFS instrument network shall be capable to handle these data flow with some burst-liked additions considered, and also is better to have capability for future upgrades of the instrument.
PFS plans to have its own network subnet with servers for DNS, DHCP, NTP. Network switches are planned to be managed under CDM (for modification of configuration, monitoring and alert handling), and also are monitored by PFS for instrument health checks during observation (e.g. to check network data flow for image data) and a data flow optimization.
This section is organized as follows:
Requirements
L1 design
  Performance verification is in next section – “Link over fiber on telescope”
  Handling on instrument exchange is in the last section – “Procedure on instrument exchange”
L3 design
Monitoring

## Slide 8: Requirements – Network
Requirements from Subaru point of view:
Subaru standard recommendation of network switches
  Cisco 2960X-24TD-L or 2960CG-8TC-L is the standard (to keep management and spare handling simple)
  If other(s) is desired, Cisco 2960X, 2960CG or 3650 are possible selection
Network switch management
  Network switch shall be managed by and monitored from Subaru CDM
  Their configuration shall follow standard configuration (e.g. VLAN, remote access) of CDM.
  One VLAN ID will be assigned for PFS-LAN
Seemless access from Subaru network (incl Hilo, VPN) to PFS LAN for remote maintenance and operation are required, although security of instrument operation need to be secured in a required level.
All connections to the Subaru dome are in fiber connection.
Static IP address assignment is desired.
Requirements from PFS instrument point of view:
Required data flow rate
  CB2F – TUE-IR (SpS): (assuming) min 200Mbps for all IR detectors (one stream), currently on TCP but alternative is on UDP
  CB2F – Cs (MCS): not assuming real time data transfer, but raw stream in ~900Mbps
  PFS to Gen2 at CB2F: not well defined but as large as possible (item in discussion session)
Special requirements on numbers of fibers:
  PF or TUE-Opt – CB2F: two pairs of Ethernet, and two pairs of single mode fiber for PCIe bus extender

## Slide 9: L1 design – Network

## Slide 11: L3 design – Network

## Slide 12: Monitoring – Network

## Slide 13: Link over fiber on telescope

## Slide 14: Storage and image data handling

## Slide 15: Requirements from ICS actors’ point of view
4 continuous stream from IR detector control computer
  Assuming 4k sq. image per 3-4sec, 60-90Mbps for one up-the-ramp
    Could be double, depends on reference pixel : TBC
    2.4TB/night for 90Mbps x 4 with 15hours operation per night, 36TB/run for 15day run
  No need to read up-the-ramp raw data
    Also saved on memory, used for final (summed?) FITS image output
VM host images
  Data rate could be small for normal operation
  NFS or iSCSI on virtio – need study from DR or backup point of view
  30GB per VM could be enough, 0.6TB for 20 VMs; 2TB with 3 generation back-up
NFS host for CCD and IR (final) FITS images
  Not high data rate – e.g. 512MB in ~40sec
  Both write and read
    Read from FTP server for sending to Gen2 (archiver)
  50 – 100 GB per night, up to 2TB for 15day run
    Survey with shortest exposure will be 7.5min (cosmology), ~100 times per night for 12 hour observation
NFS host for AG images
  1 or 2 x 2MB image per a second, both read (TBD) and write
    Focusing image will be one order smaller (e.g. one exposure per 10 sec)
  Will PFS save locally or not is a question, but ~200GB per night, ~3TB per 15day rurn
PostgreSQL data store
  Could not be on NFS, need to be direct iSCSI mount at host
  Tron log archive will be < ~1TB (SDSS 15 years in 500GB or so)
  Survey and exposure execution database will be smaller
Backup storage
  Mostly at daytime job

## Slide 16: Requirements from ICS actors’ point of view – Summary
Storage performance requirement summary
Read/Write performance
  800Mbps (90Mbps x 4SM x 2 for reference pixel + buffer) write performance at least
  Write mostly in sequence write (FITS data), ~10k IOPS is fine (or already over-spec)
Capacity – ~100TB
  10TB for operational storage (e.g. VM, DB)
  90TB for FITS-liked data; 40TB for 15day run
    Mostly raw up-the-ramp data
      Could be compressed after run -> 1/3 or 1/4
      Seeking possibility to compress at the first moment
    If compressed, 15TB for 15day run, 90TB for half year
      Will be moved to external storage in such periods
  10TB for backup of DB

## Slide 17: Hardware implementation plan
120TB (or 108TB) iSCSI RAID6 4U storage server
  20D+2P+2H x 6TB NLSAS HDD, or 2x(9D+2P)+2H NLSAS HDD
  8x 1Gbps metal interface for data, 2x metal for console interface
    device-multipath or independent at initiator
  See next for details
2 core switches for PFS network at control building
  Connected by FlexStack, might no issue on data exchange
    FlexStack has 40Gbps in total with dual link
  Connect all 10 metal to one switch could be better for management?
    Need to develop plan further on connections
  Connecting high-bandwidth lines to storage side could be possible
    Each switch has 24 metals and 4 SFPs
    All IR detector control could be connected with 4 metals and 4 SFPs
    Link to Subaru core switch could be 1 (or 2?) metals
    With them all, 16 metals and 4 SFPs in total

## Slide 18: Hardware implementation plan – CDS4004
Performance
  100K Read IOPS, 650K IOPS from cache
  6.4Gbps read, 5.3Gbps write
Host interface
  2U12 3.5in as one host or JBOD, up to 3 JBODs (+ 1 host)
    Using 1host + 1JBOD with fully populated disk (24 HDDs in total)
    Could add another 2 JBOD, if necessary during operation
  Two host controllers in redundant
    No battery backed-up cache, write from memory to flash by capacitor backed-up
  Max 8 SFP+, compatible with FC, iSCSI (hybrid) or SAS
Certifications
  NEBS 3 compliant for telecommunication
  ML-STD-810F/G
Already procured, under test at IPMU

## Slide 19: Storage coordination plan as storage server
6 LUNs for 108 or 120TB
  4LUNs for IR detector controller storage, 20TB each
    Capable to store 30day’s run
    Or 90-120 day’s if compressed
  1LUN for NFS, 25TB+
    Data, VM image storage, backup
  1LUN for PostgreSQL database, 3TB
1 NFS server will supply all required NFS/FTP storage
  If virtio IOPS requirement appears huge, might be divided into two
    If all in linux with certain version, no issue found on testing environment.
  FTP for transferring data to Gen2
    Assuming g2t can specify external FTP server by configuration
    Even we need FTP service on each, all are read only and no issue on locking.

## Slide 20: Storage coordination plan per a service
IR detector
  Direct iSCSI storage LUN mount to each control host for raw up-the-ramp FITS images
    4 iSCSI targets (LUNs) in total for this
  NFS mount for final FITS images
  1Gbps up/down full link is enough for all data I/O
    Nearly same in and out (90 or 180 Mbps) for up-the-ramp, even if no compression
VM host storage
  TBD, but assuming on NFS (qcow2?)
    Online backup is possible with qcow2 disk image
  If some writable area required, mount NFS storage at OS level (not mounting additional virtio)
  Most of hardware connected actors will run on physical machine, not VM
General NFS
  1Gbps up/down full link is enough?
  One dedicated hardware is required for NFS service
Services with special care required
  Database (PostgreSQL) storage
    iSCSI mount storage area or large volume in VM disk image (No NFS mount as filesystem!)
    Backup operation will be on pgsql layer – no special care required on storage connection
    iSCSI mount storage is better?
  Storage for backup
    Will be mostly used at daytime – daily cron job liked
    No serious performance requirement will be applicable, rely on general NFS service should be fine

## Slide 21: Plan before deployment at on-site
No special care is required, except for IR detector control
  Neither volume nor bandwidth is high
  AIT local area NFS service is enough… Or even local fs.
    From application level, nothing will differ for this use b/w iSCSI or NFS/local
      e.g. lock, bandwidth
Volume size for storage of IR detector control
  Bandwidth per one camera is not so high (~100Mbps)
    Simple NFS server can supply ~600Mbps
  Continuous up-the-ramp outputs large size data is an iusse
    100Mbps over 1day is ~1TB
    Compression will make this 250 – 350 GB/day
    But can be done with recent HDDs, without considering about backup
      No need to backup these software stability test data
  Need to be studied once PFS get an IR detector
    e.g. how much we need to save/check/analyze up-the-ramp data
    Better to study how existing instrument at Subaru doing?

## Slide 22: Emergency recovery plan
(to be studied following failure modes)

## Slide 23: VM infrastructure

## Slide 24: Database server and its replication/backup

## Slide 25: Status monitoring and defect detection

## Slide 26: Procedure on instrument exchange

## Slide 27: Hardware delivery to Subaru

## Slide 28: Items and assigned delivery
Followings are hardware items which the project (or Subaru) need to prepare or deliver to the summit, and subsystems use these items.
CB2F (Control Building 2F)
  Rack to host items (Subaru): before all of followings
    Assumed deepest unit is ~760mm (MPS), ~650mm (DB host)
  Network switch (IPMU; already in hand): (at least) 1 at MCS, 2nd at SCR
  ICS RAID6 storage (IPMU; already in hand): SM1
    MCS backup could be done with small storage?
  VM host computers (IPMU; some already in hand)
    (at least) 1 at MCS (any subsystem delivered to summit) – to run MHS
    Others depends on system load
  Serial communication (MLP1; IPMU) host: PFI
  KVM, PDU etc. (IPMU; some already in hand): MCS?
  DB host (IPMU; spec TBD): SM1??
SpS (both IR3 and TUE-IR)
  5th rack (from JHU): SCR
  Network switch to connect CB2F (IPMU; already in hand): SCR
  UPS and power supply (Subaru): SCR
  1U host computer @ 5th rack (IPMU; spec TBD): SCR
  Telemetry devices/system (TBD): SCR (or SM1)

## Slide 29: Setup to accept MCS
MCS will be installed as Cs instrument. We need to have minimum operation hardware at CB2F, including PFS network.
19inch Rack (Subaru)
PFS network switch (at least 1; IPMU; already in hand)
  Also connection to subaru’s core switch
VM host (IPMU; 1 already in hand)
  Hosting administrative VM, e.g. dnsmasq
KVM, PDU (IPMU; already in hand)
At this stage, no special care for power distribution is required, only like connecting things to PDU at CB2F. Also network is simple, we need two fiber connection to Cs and its standby (one active, one backup), one connection to subaru core switch with routing.

## Slide 30: Setup to accept SCR
In parallel to build SCR (Spectrograph Clean Room) at TUE-IR (IR4), we need to have control hardware for SCR.
2nd network switch at CB2F (IPMU; already in hand)
  might only need 1 SFP to SpS, but better to have and check working at this moment
Network switch at SpS (IPMU; already in hand)
5th rack (from JHU)
UPS at IR3, with power distribution/panel (Subaru)
1U host computer (also for SCR control) (IPMU)
Telemetry devices (TBD)
  e.g. Lakeshore temperature monitor and sensors
At this point, we need full power distribution system to supply spectrograph, both UPS and power panel w/ cables. Environment active control subsystem for SCR shall be considered as a part of SCR, so not listed here.
For network and storage, we only need “connected” at this point, so performance evaluation on network in parallel to SCR preparation (or validation) is fine.

## Slide 31: Setup to accept SM1,2
Before we have SM1,2 at summit, we need almost all of hardware.
ICS RAID6 storage (IPMU)
DB host (IPMU)
  Not as full operation (e.g. interaction with ETS or survey coordination), but to store MHS status archive or telemetry etc.
Before this point, PFS network and storage need to be fully prepared and validated on performance point of view (for NCU1,2 operation).

## Slide 32: Setup to accept PFI
Most of computing and power resources need to be active well before delivery of PFI to the summit.
Serial communication (IPMU)
  This communication will be validated at factory of melco (Mitsubishi; at Amagasaki) in test setup.

## Slide 33: Discussions on interfaces between Subaru and PFS
Network connection to Subaru (data flow, V-LAN)
Access control to/from PFS network
System alert handling (items to STS)
Engineering data archive (during on-site AIT, independent unit test)

## Slide 34
Backup or copied materials to be used

## Slide 35: Updates/Progresses – shared infrastructure
Communication fibers at Subaru have tested and passed
  Mostly by real data flow, except for 1Gbps SM fibers (OTDR)
ICS storage performance have verified with simulated 4 H4RG data streams
  Simulation configuration: Storage – NFS server – 4x data generator
  Achieved ~2.3Gbps file write, enough for 4 H4RG data streams
  ToDo
    Need some consideration on network configuration
    More realistic test (e.g. having both read/write) could be better
PFS-LAN network host stage for configuration and stability passed
  Using real hardware to be used on-site
  Some management configuration will be modified by Subaru (CDM) after delivery
  ToDo
    Connection between PFS-LAN and Subaru network is not fixed yet
Data backup procedures (e.g. DB) are in production at pfs servers
Virtual machine servers for ICS components are under host stage at IPMU
  Dell R710 (2U Xeon E5506) x2, R410 (1U Xeon E5640x2) x3, R310 x1+ backup (2U x2)
  4 port Gbps NIC

## Slide 36
Cisco 2960X-24TS-L
fiber
metal
Cisco 2960X-24TS-L
Cisco 2960X-24TS-L
SpS SM1 switch
SpS SM2 switch
SpS SM3 switch
SpS SM4 switch
SpS (TUE-IR)
CB2F
Subaru core
JHU
IPMU
Subaru
PFI
ASIAA
ICS Storage (iSCSI or NFS)
Media Conv.
Cal Controllers
Media Conv.
PFI Controllers
Cs
2960CG-8TL-L
MPS
AG Ctrl
USB-Opt Conv. (x2)
COBRA
USB-Opt Conv. (x2)
AG Cam (x6)
POpt2 Cable Wrapper
PFS Network Physical Wiring Diagram
1 or 2 TBC??
4SM in LACP
MM 62.5
MM 62.5
MM 50
TBD
MM 50
FlexStack x2 (dual link)
Cs Stand-by
8x 1Gbps metal in device multipath iSCSI
Exchanged at Cs only
Cs focus
Different fibers at PF exchange
Inst Exchange
Connected
all time @CB2F
E-LAN

## Slide 37: Network connection ToDo
Connection to Summit core (E-LAN, V-LAN)
  To send up-the-ramp smoothly, we may need more than 1Gbps. Need LACP configuration.
    PFS side for up-the-ramp to Gen2 is one g2cam host (TBC), need some care on LACP balancing.
  V-LAN is also through Summit core, could connect through trunk port configuration on Summit core and PFS core.
    Even if we select this, I want to keep separate connection for V-LAN and PFS-LAN to the host (in PFS network).
Connection between PFS core and SpS (ICS #579)
  There is a way to go on configuration to pass 4x300-400Mbps stream, but need to be cared in configuration both BEE and iSCSI storage.

## Slide 38: Host cluster at CB2F
Most of services are planned to run on VMs, such as MHS, logger, support infrastructure. Also PFI related actors are in consideration.
Some services (esp. DB) is planned as bare system, but could be on VM if performance is verified.
Host configuration (still in development):
iSCSI storage (108TB RAID6 w/2sp, 2x2U) x1
Dell R710 Xeon E5506 (2xAC, 2U) x2
Dell R410 Xeon E5640 (1xAC, 1U) x3
Dell R310 Xeon E3-? (1xAC +1 vacant, 1U) x1 (for DB?)

## Slide 39: Host cluster ToDo
Verify iSCSI storage performance with more realistic configuration
  Currently only write from 4xH4RG simulator – NFS server – iSCSI storage.
  Better to verify with running VMs, DB connections, with some data read/write (incl. sending up-the-ramp to Gen2)
IR data handling
  Original plan was host control computer at CB2F with large memory, which will do all for data management
    Read directly from SAM over 1Gbps network
    Save up-the-ramp FITS (single 2D HDU) to storage, and on memory
    Build final FITS file from data on memory (no read stream from storage for this)
  Currently, BEE reads data from SAM via USB, and saves to storage
    Something other need to care of building final FITS file etc.
Verify VM operation performance
  Mostly tested using PFS server cluster at IPMU for pfs.ipmu.jp
Integration of PFI control into VM cluster
  Need to check PCI passthrough performance and capability for PCIe extender.
    Or USB mount could work?
  Check network bandwidth requirement (for shared interface)
Host maintenance operation
  Health check with iDRAC etc.