## Slide 1: Procedure to measure and store MPS configurations
Atsushi Shimono
(2015/11/27)
PFS Software team Study Note : SSN-00021-001+

## Slide 2: Outline
Aim of this note is to summarize required procedures to measure and store MPS (Cobra) operational configurations. Parameters of Cobras, such as arm lengths and limit angles, will be measured using MCS at the summit, and will be stored into MPS database. (Also measured with some camera at ASIAA during AIT; but need to be re-measured at the summit.) All these measurement and configuration will be written in F3C coordinate (refer commissioning and operation work flow), conversion to F3C from measured values are not included.
This will not include configurations of motor calibration data, like operational frequency (of piezo?).
Required subsystems and their setup
Procedures to measure operational configurations
Configuration data and revision handling
ToDo

## Slide 3: Required subsystems and their setup
Baseline concepts of operation for measurements are to get shots by MCS with moving Cobras in various motions.
Required operational modes of subsystems are:
FPS, MPS: need special operation mode/sequence
SpS/fiber back illumination system: strobe mode
Fixed fiber back illumination system: normal operation
MCS: longer exposure than normal operation

## Slide 4: Procedures to measure operational configurations
PFS will not have any precise-enough mechanical measurement of positions (and configurations) before taking image by a camera, like MCS (how to do during AIT is another question, though). This procedure need to be capable measuring all configuration parameters required to define movements of COBRAs. Also, at every step, need to be safe on collision avoidance point of view.
Considering mechanical design and precision on fabrication, we can assume followings:
Center of COBRA patrol area is nearby (~1%) vertex of triangles
Length of each stage (arm length) is in ~1% accuracy
2nd stage rotation limits (CW/CCW) are close to direction of 1st stage arm
ToC are:
Configurations to be measured
Operations

## Slide 5: Procedures (I): configurations to be measured
Position
  (x, y) of center at F3C
Length
  Theta (1st) stage
  Phi (2nd) stage
CW/CCW limit angle
  Theta stage to F3C
  Phi stage to Theta origin
    Designed to close at 1st arm direction
Size (how to measure?)
  2nd stage size (outer of arm)
  Arm size (from center of 2nd stage)
Closer limit to center

## Slide 6: Procedure (II): Operation (1) – center (x,y)
First, measure center positions of COBRAs, by rotating only 1st stage and measure the center of dotted arc made by strobe mode.
Before moving 1st stage, take spots without moving positioner motors, to check possible collision before moving anything. Using center positions and stage lengths of mechanical design, the farthest points from each center position can be calculated in a few % accuracy.
If some COBRAs are marked as possible collision, move 2nd stage of target (one by one) to direction to inner side (which would be defined by mechanical design).
Once ready to measure:
Start MCS image acquisition with strobe mode and moving only 1st stage
  Move 1st stage from end to end (CW limit to CCW limit)
Dotted circle image will be acquired, on F3C
Measure dots of dotted circle, and calculate center of circle
These points are (x,y) of center positions of COBRAs on F3C.

## Slide 7: Procedure (II): Operation (2) – 2nd stage
To measure arm length of 2nd stage, dotted circle similar to operation (1) but with moving only 2nd stage motor is in need. With relying on mechanical design, angles to move microlens close to center position are fixed (as direction – CW or CCW), and it should be safe to move 2nd stage 45 degree from the closest limit to center position, which makes all of 2nd stage arm well within non-collidable area.
Operation:
Start MCS image acquisition with strobe mode and moving only 2nd stage
  2nd stage from CW or CCW limit (the closest to center position) with 45 degree
Dotted circle image will be acquired, on F3C
Measure dots of dotted circle, and calculate circle diameter
Results of circle diameter are arm lengths of 2nd stage.
With this safer operation, we can only take up to 60 degree (TBC), which is only 1/6 of full circle, and results could be somehow inaccurate….

## Slide 8: Procedure (II): Operation (3) – 1st stage
Now, we have center positions (x,y) and arm length of 2nd stage.
Using arm length of 2nd stage, we can evaluate (somehow??) accurate delta angle between two configurations of 2nd stage. As in next slide, measuring with three angles of 2nd stage, we can calculate arm length of 1st stage.
Operation:
Move 2nd stage in safe region, and do the same for two other angle
Measure dots of dotted circles, calculate arm length of 1st stage
From operation 1 to 3, we now have center positions (x,y) and two arm lengths. Assuming 2nd stage rotatable angle within 0 to 180 degree from 1st stage direction, we can convert position (x,y) on F3C to (theta, phi).

## Slide 9
A1
A2
L_phi
L_theta
phi_1
phi_2
A1^2 = L_phi^2 + L_theta^2 – 2L_phi*L_theta cos(phi_1)
A2^2 = L_phi^2 + L_theta^2 – 2L_phi*L_theta cos(phi_2)
->
A1^2-A2^2=4L_phi*L_theta sin((phi_1+phi_2)/2)sin((phi_1-phi_2)/2)
Since phi_1-phi_2 is measured as a difference, A_1 and A_2 are measured as diameter of dotted circles moving 1st stage, and arm length of 2nd stage (L_phi) is already measured, this equation only has parameters of L_phi and phi_1 (or phi_2).
So, if we measure three different angles of phi, we can solve this equation to get arm length of 1st stage.

## Slide 10: Procedure (II): Operation (4) – CW/CCW limits
From operations 1 to 3, we have center positions (x,y) and two arm lengths. Remaining values are CW/CCW limits of rotation angles for both 1st and 2nd stage.
These limits will be easily measured. To make possibility of collision as small as possible, we are better to measure from 1st stage limits and “closer limit to center” of 2nd stage – since rotating 1st stage with setting 2nd stage at other limit of 2nd stage will make arm of 2nd stage collide.
Operation:
Move stages at limits to be measured
Take normal MCS image, with normal illumination (w/o strobe)
Calculate dot positions and convert into angles

## Slide 11: Configuration data and revision handling
Configuration data will be put into database of MPS via API call through FPS. Also MPS has functions to import/export via XML format.
We can save the XML file into some repository with version ID, which is the identification.

## Slide 12: ToDo
Evaluate accuracy we can acquire from each operation 1-4
Define format of XML data file